## Slide 1
Status and R&D Progress of
Detector Magnet
Zian Zhu
Institute of High Energy Physics
(zhuza@ihep.ac.cn)

## Slide 2: Outline
Operating status of BESIII magnet
R&D related to detector magnet at IHEP
Summary

## Slide 3
BESIII magnet
BESIII is the third upgrade of Beijing Spectrometer that run with BEPC
1 T, length 4 m, bore diameter 2.7 m

## Slide 4: Coil winding & magnet installation
Coil winding
Assembly
Installation（2006.9）
Turn magnet from H to V

## Slide 5: Magnet current
Histogram of current during a running period

## Slide 6: Quenches of BESIII magnet
NO. | Time | cause of quench | Quench current
1 | 2008.7.29 | SCQ’s quench | 3034A(0.9T)
2 | 2008.8.18 | SCQ’s quench | 3369A(1.0t)
3 | 2008.9.23 | Current ramping down too fast | 3190A
4 | 2008.12.18 | SCQ’s quench | 3369A
5 | 2009.2.26 | Unknown | 3369A
6 | 2009.3.7 | Unknown | 3369A
7 | 2009.3.27 | Power grid fault | 3369A
8 | 2009.4.25 | Unknown | 3369A
9 | 2009.5.7 | Cryogenic system failure | 3369A
10 | 2009.5.31 | Cryogenic system failure | 3369A
11 | 2009.12.25 | Cryogenic system failure | 3369A
12 | 2010.1.13 | Vacuum system failure | 3369A
13 | 2010.11.16 | SCQ’s quench | 3369A
14 | 2011.1.13 | Cryogenic system failure | 3369A
15 | 2011.1.17 | Cryogenic system failure | 3369A
16 | 2011.5.10 | Cryogenic system failure | 3369A
17 | 2011.5.26 | Cryogenic system failure | 3369A
18 | 2011.6.1 | Cryogenic system failure | 3369A
19 | 2012.1.29 | Cryogenic system failure | 3369A
20 | 2012.2.3 | Cryogenic system failure | 3369A
21 | 2012.2.14 | SCQ’s quench | 3369A
22 | 2012.3.23 | Power grid fault | 3369A
23 | 2012.11.15 | Power grid fault | 3369A
24 | 2013.2.10 | Cryogenic system failure | 3369A
25 | 2014.03.12 | Cryogenic system failure | 3369A
26 | 2014.05.26 | Quench detector misoperation | 3369A
27 | 2015.02.20 | Cryogenic system failure | 3369A
28 | 2015.05.18 | Power network problem | 3369A
29 | 2015.12.30 | Quench detector misoperation | 3200A
30 | 2016.01.17 | Cryogenic system failure | 3369A
31 | 2016.02.9 | Cryogenic system failure | 3369A
32 | 2017.04.13 | Power Fluctuation | 3369A
33 | 2017.05.07 | Cryogenic system failure | 3369A
34 | 2017.11.28 | Cryogenic system failure | 3369A
35 | 2018.03.05 | Cryogenic system failure | 3369A
Total 35 quenches
18 cryogenic system
5 electricity power
5 SCQ quench
2 quench detector
1 vacuum
1 operation error
3 unknown

## Slide 7: Aging Problem－－Vacuum
From 2010 to now, the vacuum value kept rising from 2.2×10-2Pa to 2.7×10-2Pa, even to 3.8×10-2Pa in 2016,and back to 2.7×10-2Pa ,it was not stable within last few years
It increased significantly when temperature is under 20K, it should be low temperature leakage

## Slide 8: 2017-1-8 23:50，Vacuum and pressure of Dewar jumping
The pressure of the Dewar is up to 1.41bara (normal is about 1.36bara)
The Dewar liquid level is normal
Insulation vacuum and small Dewar pressure jumped simultaneously , they become to normal after 40 hours.

## Slide 9: Aging Problem－－Current Transition section in valve box
From Feb.2010, the temperature of the transition section(TSD12) kept rising to 6.42K

## Slide 10
Aging Problem－－Transition section in valve box

## Slide 11
We have done temperature margin test in Nov.12 2012, and adjusted warning threshold from 6.38K to 6.5K(TSD-12, temperature of transition section)
We worry about :
1.Temperature keep rising and out of control
2.Burn the current transition section
Aging Problem－－Current transition section in valve box

## Slide 12
Lots of cryogenic pipelines and cables at the top of the valve box.
Small distance between valve box with other devices at three sides, and resulting in inconvenient assembly.
New valve box
Aging problems become more and more obvious as running time increased
Replacement time is expected within two months with a backup of valve box . Otherwise it maybe take one year to solve the accident
Surroundings of Valve Box

## Slide 13: Requirements and technical improvements
Requirements：
Unchange the overall dimensions of the valve box
Unchange the position and mode of the joint with the cryogenic system
Unchange position of the joints with power bus
Overall heat leakage can not be increased
Technical improvements:
HTS current leads instead of gas-cooled current lead
Optimize the cryogenic pipeline, reduce number of the bimetal joints, to improve the vacuum
Optimization of cooling structure of the transition section
Inside structure of the existing valve box
Design of new the valve box

## Slide 14: New current leads
Operating current: 3400A
Structure: binary current leads
Heat exchange section cooled by helium gas
HTS stack structure

## Slide 15: Cryogenic pipeline
Pipes from the coils are all aluminum tubes
There are several bimetal joints in the valve box, and these maybe the reason of the bad vacuum
Optimize the cryogenic pipeline:
1. Reduce number of the bimetal joints
2.Remove the cooling stations
Bimetal joints
Insulators
To improve the vacuum

## Slide 16: R&D related to detector magnet at IHEP
Requirements of next generation detector:
a larger size and higher field magnet
Strong Solenoid
  Better Momentum resolution for energetic tracks
  Better separations
Baseline Value: 3 Tesla, 7m bore diameter
This R&D work is supported by Innovation Fund of IHEP for CEPC

## Slide 17: Starting point: BES-III Magnet
Aluminum
Mechanic support
Nb-Ti Superconductor
(Rutherford Cable)
Embed in pure Aluminum
Structure
Pipe for
Liquid He at 4.2 K
Cutting slice from the BESIII magnet prototype

## Slide 18: Special Cable
From BES → CEPC detector
Current: 4 kA → 16 kA
3-4 layers of cables

## Slide 19: Rutherford Cable - I

## Slide 20: Rutherford Cable - II

## Slide 21: Rutherford Cable - III
RRR: Ratio of Residual Resistance
Ic: Critical Current

## Slide 22: Extrusion of Aluminum - I

## Slide 23: Cladding of Aluminum

## Slide 24: Inserting Process

## Slide 25: Shear Strength

## Slide 26: Magnet Cooling

## Slide 27: Magnet Cooling Prototype

## Slide 28: New idea I: High Temperature Superconductor

## Slide 29: Summary
10 years stable running of BESIII detector magnet
  Prepare to make a new valve box to avoid potential risk of long period machine shut down
Recently R&D milestones
  Constructed 10 meter long aluminum stabilized NiTi Rutherford superconductor of 10kA@4T
  Key performances: RRR, Ic, Shear Strength meet the requirement
To do
  >100 meter long superconducting cable development
  Magnet study under high radiation dose
  Thermal-siphon liquid helium cooling system study
  HTS detector magnet study

## Slide 30: Thank you